## Slide 1: TES Safety and Tolerability
Dr. Michael Amlung
12/8/2017

## Slide 2: Disclaimer
The content presented in these slides is meant to provide an overview of current best practices and guidelines for safety of TES. These recommendations are not intended to be exhaustive and may become outdated as additional research on safety is published.
The author does not guarantee that this content is up-to-date if these slides are accessed at a later date.
Always consult current safety guidelines before initiating any TES research or clinical project.

## Slide 3: Topics
Is TES non-invasive?
Key terminology & definitions
Adverse effects of TES
Monitoring adverse effects after TES
Contraindications for TES
Recommendations & conclusions

## Slide 4: Is TES Non-Invasive?

## Slide 5: Consensus Group Paper

## Slide 6: Current Comprehensive Review

## Slide 7: Key Terminology & Definitions

## Slide 8: Key Terminology & Definitions
Tolerability
  Presence of uncomfortable and unintended effects, which do not induce structural or functional damage
Safety
  Refers to occurrence of damaging effects
  Can only be considered in relative terms
    Source: Brunoni Chapter 12 Transcranial Direct Current Stimulation in Neuropsychiatric Disorders

## Slide 9: Key Terminology & Definitions
Side Effects vs. Adverse Effects
  Side effect and adverse effects often used interchangeably, but this is problematic
  According to Antal et al (2017), Side effects are consequences that are different than the intended effect, and might be good or bad (beneficial or adverse).
  Side effect example: Improvement in memory by an intervention for depression
  Adverse events are, by definition, always bad.

## Slide 10: Key Terminology & Definitions
Adverse Effects
  Food and Drug Administration (FDA): Any undesirable experience associated with the use of a medical product (or intervention) in a patient or participant
  May or may not be causally related to the intervention
  Mild AE: Mild symptoms for which no medical treatment is necessary (skin redness, tingling)
  Moderate AE: Need of local or noninvasive treatment (application of cream after a skin burn)
  Serious AE: Severe or medically significant but not immediately life-threatening events, include the inpatient hospitalization
  Life-threatening SAEs: Any event that is life threatening or resulting in death.

## Slide 11: Key Terminology & Definitions
Adverse Reaction (Health Canada)
  Undesirable effects to health products, including drugs, medical devices, and natural health products. Reactions may occur under normal use conditions of the product.
  Reactions may be evident within minutes or years after exposure to the product and may range from minor reactions like a skin rash to serious and life-threatening events such as a heart attack or liver damage.
  Governed by the Food and Drug Act and Regulations

## Slide 12: Key Terminology & Definitions
Good Clinical Practice (GCP)
  Adverse Event (AE)
  Any untoward medical occurrence in a patient or clinical investigation subject administered a pharmaceutical product and which does not necessarily have a causal relationship with this treatment. An adverse event (AE) can therefore be any unfavourable and unintended sign (including an abnormal laboratory finding), symptom, or disease temporally associated with the use of a medicinal (investigational) product, whether or not related to the medicinal (investigational) product.

## Slide 13: Key Terminology & Definitions
Good Clinical Practice (GCP)
  Serious Adverse Event (SAE)
Any untoward medical occurrence that at any dose: results in death, is life-threatening, requires inpatient hospitalization or prolongation of existing hospitalization, results in persistent or significant disability/incapacity, or is a congenital anomaly/birth defect

## Slide 14: Key Terminology & Definitions
According to European Medical Device Directive
  “Safe” is a condition where all risks are accepted risks
  However, all stimulation protocols carry a certain degree of risk and could cause problems in specific circumstances
  Current approach is to estimate potential of a protocol becoming a hazard, or a potential for an adverse event
Risk is a measure of the combination of the hazard, the likelihood of occurrence of the AE and the severity.
Risk Benefit Ratio: Overall ratio of all potential benefits of a procedure divided by all the adverse reactions.
Antal et al 2017

## Slide 15: Key Terminology & Definitions
Direct benefit, high risk: Diagnostic or therapeutic primary objective, direct clinical benefit for participant. Acceptable risks may be higher. Generally patients only.
Indirect benefit, moderate risk: Little or no expectation of clinical benefit. Development of treatments, safety, understanding diseases. Generally patients only (unless healthy controls are included as a comparison)
Indirect benefit, low risk: Important data on brain physiology, general pathology, or safety, but no immediate clinical relevance
High benefit, low risk: Important data on cognition/brain function, immediate relevance for cognitive/motor improvement (neuroenhancement)
Antal et al 2017

## Slide 16: Key Terminology & Definitions
Risk is differentiated from burden: a procedure may be burdensome (e.g., produce much discomfort) but still be considered safe (e.g., not having any relevant risk for permanent damage)
Antal et al 2017

## Slide 17: Adverse Effects of TES

## Slide 18: Start with the Take Home Message
“Within the standard parameters of use… the evidence indicates that tDCS is a well-tolerated technique, with few, mild side effects”
“Although tDCS is considered to be ‘safe’, as the tDCS device is limited to delivering a low-dose current which has effects on cortical excitability, and no major or serious adverse effects have been reported, such findings do not imply that tDCS is ‘universally safe’ and should therefore be used without limits or controls”
Brunoni et al. Chapter 12. p 348

## Slide 19: Tolerability: Common Adverse Effects
Poreisz et al. 2007 – Data from 567 tDCS sessions
  Mild tingling sensation (71%)
  Fatigue (35%)
  Itching (30%)
  Headache (12%)
  Nausea (3%)
  Insomnia (<1%)

## Slide 20: Tolerability: Common Adverse Effects
Systematic reviews, meta-analyses, and other studies
  Poreisz et al. 2007 – Data from 567 tDCS sessions
    Mild tingling sensation (71%)
    Fatigue (35%)
    Itching (30%)
    Headache (12%)
    Nausea (3%)
    Insomnia (<1%)
All effects were mild, short-lived, and well-tolerated. Rate was mostly similar for active/sham conditions

## Slide 21: Common Adverse Effects
Systematic reviews, meta-analyses, and other studies
  Brunoni et al. 2011 – Meta-analysis of all tDCS clinical studies from 1998 to 2010.
  209 studies, nearly 4000 subjects
  56% monitored affects, 63% reported at least one averse effect

## Slide 22: Common Adverse Effects
Systematic reviews, meta-analyses, and other studies
  Brunoni et al. 2011 – Meta-analysis of all tDCS clinical studies from 1998 to 2010.

## Slide 23: Common Adverse Effects
Skin Reddening: tDCS-induced erythema
Intensity varies in patients; most only experience mild redness, while others might have more intense reddening
Cause is not clear, but likely partially due to increased blood floor in dermal vessels; release of neuropeptides, vasoactive substances
More prominent over anode than cathode
Lasts less than 18-24 min
Less intense in darker skin color
Not influenced by gender, age, or smoking status

## Slide 24: Common Adverse Effects
Contact dermatitis:

## Slide 25: Serious Adverse Effects
“To date, the use of conventional tDCS protocols in human trials (≤40 min, ≤4 mA) has not produced any reports of a Serious Adverse Effect or irreversible injury across over 33,200 sessions and 1,000 subjects with repeated sessions.
This includes a wide variety of subjects, including persons from potentially vulnerable populations.”
Source: Bikson et al. (2016). Safety of transcranial direct current stimulation: Evidence based update 2016. Brain Stimulation

## Slide 26: Serious Adverse Effects
Includes induction of seizure, stroke, cardiac arrest, etc.
  One report seizure in a 4-year-old with history of seizure/spasm
  Screen for history of seizure, epilepsy, etc.
Safety studies indicate:
  tDCS does not change heart rate variability
  Does not increase level of brain enzymes linked to neuronal death
  Does not qualitatively alter EEG activity

## Slide 27: Serious Adverse Effects
Skin Lesions/Burns
  Palm et al. 2008 reported 5 cases of skin lesions after 5 days of 2mA stimulation with tap water soaked sponges
  Frank et al 2010 reported 3 cases of skin lesions under anode 1.5ma, tap water soaked
  Rodriguez et al 2014 reported 4 cases of skin burn under cathode, saline soaked

## Slide 28: Serious Adverse Effects
Skin Lesions/Burns
Wang et al. 2014 Brain Stimulation Letter to Editor.
5x7cm electrode, saline soaked, cathode secured by elastic strap, 2mA stimulation for 26 min.
Subject retrospectively reported a burning/stinging sensation during tDCS
“We believe that this skin burn resulted from extensive skin heating produced by inhomogeneous electric current which resulted in higher temperatures in a local area, the middle area under the electrode in the forehead region.”
Conclusion  Use caps, not straps! And immediately report any discomfort

## Slide 29: Serious Adverse Effects
Skin Lesions/Burns
  Brunoni chapter conclusion: Skin damage caused by tDCS has been occasionally reported. Unclear whether it is more common under anode or cathode. More common with tap water soaked sponges.
  Loo et al. 2011 suggest some precautions
    Screening for skin disease
    Checking skin where electrode is placed before each session
    Avoid abrasion of the skin
    Ask patients to report any pain during tDCS (but not foolproof  Palm et al 2014)

## Slide 30: Serious Adverse Effects
Neuronal Damage / Lesions
Animal study by Liebetanz et al (2009)

## Slide 31: DIY TDCS
Bikson et al (2013) quote on DIY:
“Controlled investigation of tDCS for treating neuropsychiatric disorders or for neurorehabilitation should not be confused with improvised devices...
Experimentation outside established and tested norms may put subjects at risk…
Meddling with the tDCS dose is potentially as dangerous as tampering with a drug’s chemical composition.”

## Slide 32: Special Populations
Children
  See section in Bikson et al (2016) Brain Stimulation review
  See section in Antal et al. (2017) Clinical Neurophysiology consensus paper
  “In trials involving children, at least 2800 sessions have been applied across nearly 500 subjects. No serious adverse effects have been reported. tDCS has been investigated in children with a variety of diagnoses including cerebral palsy, stroke, encephalitis, epilepsy, schizophrenia and attention-deficit hyperactivity disorder.” Bikson 2016
  Possible increased brain current density at same voltage (smaller average head size)
  “Combining this emerging pediatric evidence with the larger animal and adult experience suggests tDCS within the same ranges of dosing and duration can be considered minimal risk, based on current evidence, in school-aged children.” Bikson 2016

## Slide 33: Special Populations
Aging
  See section in Bikson et al (2016) Brain Stimulation review
Mood Disorder
  Induction of (hypo)manic episodes.
  There are four stand-alone case reports in literature and some reports in randomized clinical trials of mania or hypomania induction after tDCS treatment.
  Difficult to infer whether tDCS caused symptoms or if they were a part of natural history of the disease
  Bikson et al (2016) “emphasize a lack of convincing evidence that tDCS increases risk of manic switch. In trials for depression, greater than 4160 sessions have been applied across more than 430 subjects without a documented Serious Adverse Effect.”

## Slide 34: Monitoring Adverse Effects
Consensus group provided a questionnaire for systematically monitoring AEs during stimulation protocols
Available here: http://www.neurologie.uni-goettingen.de/downloads.html

## Slide 35: Monitoring Adverse Effects
Consensus group provides a questionnaire for systematically monitoring AEs during stimulation protocols
Available here: http://www.neurologie.uni-goettingen.de/downloads.html

## Slide 36: Contraindications for TES
Not all of these are definite rule-outs: *Do not rely on this list*
  Do not place electrodes directly above areas of impaired skin
  No chronic skin diseases
  No implanted metallic plates or devices, neurostimulators, pacemakers, pumps, etc.
  No history of previous neurosurgical procedures, neurologic malformations, or brain neoplasias
  Head trauma / loss of consciousness / skull defects
  Pregnancy
  Seizures/epilepsy
  Fainting spells, syncope
  Previous adverse reactions / effects to TES or TMS
  Medications?
  Caffeine? Alcohol?

## Slide 37: Screening Participants

## Slide 38: Bottom Line
Always screen carefully
Always inspect the equipment prior to each use
  Old/dried sponges, broken electrodes, frayed wires, rusting, device charged, appropriate blinding in place, etc.
Always monitor for AEs during experiment and also consider contacting for follow-up monitoring
Always document and report any AEs
Always operate within standard protocols and procedures
Have emergency procedures in place
Training is key!

## Slide 39: Comment on Standardizing Protocols
Standard system for multidimensional parameter space for the experiment. Clearly defined protocols with electrode type, positions, current type, intensity, duration, sequencing, spacing, etc.
Use checklists provided by the Consensus Group (Antal et al. 2017)

## Slide 43: Return to the Take Home Message
“Within the standard parameters of use… the evidence indicates that tDCS is a well-tolerated technique, with few, mild side effects”
“Although tDCS is considered to be ‘safe’, as the tDCS device is limited to delivering a low-dose current which has effects on cortical excitability, and no major or serious adverse effects have been reported, such findings do not imply that tDCS is ‘universally safe’ and should therefore be used without limits or controls”
Brunoni et al. Chapter 12. p 348